## custom volume UPM034A
1 | Instrument | Brand Pipettor (code) |  | FINPIPETTE
|  | Serial Number |  | OZ58589
|  | Nominal Volume (ul) |  | 1000
2 | Physical condition | □ Good working condition □ Functional damage : leaking; display not working, button defective, battery defective.
3 | Room condition | Suhu bilik (20 ± 3°C) |  | 22.4
|  | Relative humidity (45 - 80%) |  | 65
4 | Water temperature (˚C) | 22.5
| Balance | METLER TOLEDO XP 205 DR (006-004852)
| Thermometer | THERMO FISHER LOLLIPOP (150736002)
| Z Factor | 1.0034
5 (a) | Result of gravimetric test for selected test volume (ul): |  |  | 1000
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 995.42 |  | 998.8044
| X2 | 996.52 |  | 999.9082
| X3 | 996.92 |  | 1000.3095
| X4 | 997.88 |  | 1001.2728
| X5 | 997.39 |  | 1000.7811
| X6 | 995.05 |  | 998.4332
| X7 | 995.96 |  | 999.3463
| X8 | 997.55 |  | 1000.9417
| X9 | 996.47 |  | 999.858
| X10 | 998.03 |  | 1001.4233
| Mean value | 996.719 |  | 1000.1078
(b) | Evaluation of gravimetric test
| Mean volume (Vi) | 1000.1078
| Standard deviation (sr) | 1.0254
| % A found | 0.010784
| % CV found | 0.10253
| Accuracy Limit (A%) | 0.8
| Coefficient Variation Limit (CV%) | 0.3
| Result | PASS
6 (a) | Result of gravimetric test for selected test volume (ul): |  |  | 500
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 499.01 |  | 500.7066
| X2 | 497.93 |  | 499.623
| X3 | 499.82 |  | 501.5194
| X4 | 497.99 |  | 499.6832
| X5 | 498.51 |  | 500.2049
| X6 | 499.13 |  | 500.827
| X7 | 499.74 |  | 501.4391
| X8 | 498.72 |  | 500.4156
| X9 | 499.46 |  | 501.1582
| X10 | 498.22 |  | 499.9139
| Mean value | 498.853 |  | 500.5491
6 (c) | Evaluation of gravimetric test
| Mean volume (Vi) | 500.5491
| Standard deviation (sd) | 0.69553
| % A found | 0.10982
| % CV found | 0.13895
| Accuracy Limit (A%) | 1.6
| Coefficient Variation Limit (CV%) | 0.6
| Result | PASS
7 (a) | Result of gravimetric test for selected test volume(ul): |  |  | 100
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 98.31 |  | 98.6443
| X2 | 98.97 |  | 99.3065
| X3 | 99.38 |  | 99.7179
| X4 | 98.08 |  | 98.4135
| X5 | 98.63 |  | 98.9653
| X6 | 98.06 |  | 98.3934
| X7 | 98.1 |  | 98.4335
| X8 | 98.1 |  | 98.4335
| X9 | 98.93 |  | 99.2664
| X10 | 98.94 |  | 99.2764
| Mean value | 98.55 |  | 98.8851
7 ( b) | Evaluation of gravimetric test
| Mean volume (Vi) | 98.8851
| Standard deviation (sd) | 0.48374
| % A found | 1.1149
| % CV found | 0.4892
| Accuracy Limit (A%) | 8
| Coefficient Variation Limit (CV%) | 3
| Result | PASS
8 | Interpretation | Criteria : % A found equal or not more than Accuracy Limit (A%) a% CV found equal or not more than Coefficient Variation Limit (CV%)
Penganalisa |  |  | Penyemak | Pelulus
Tandatangan : Tarikh ujian : Nama : |  |  | Tandatangan : Tarikh semak : Nama : | Tandatangan : Tarikh disahkan : Nama :

## < 50ul UPM034B
1 | Instrument | Brand Pipettor (code) |  | FINPIPETTE
Serial Number
Nominal Volume (ul)
2 | Physical condition | □ Good working condition □ Functional damage : leaking; display not working, button defective, battery defective.
3 | Room condition | Suhu bilik (20 ± 3°C)
Relative humidity (45 - 80%)
4 | Water temperature (˚C)
| Balance | METLER TOLEDO XP 205 DR (006-004852)
| Thermometer | THERMO FISHER LOLLIPOP (150736002)
Z Factor
5 (a) | Result of gravimetric test for selected test volume (ul):
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 |  |  | #VALUE!
| X2 |  |  | #VALUE!
| X3 |  |  | #VALUE!
| X4 |  |  | #VALUE!
| X5 |  |  | #VALUE!
| X6 |  |  | #VALUE!
| X7 |  |  | #VALUE!
| X8 |  |  | #VALUE!
| X9 |  |  | #VALUE!
| X10 |  |  | #VALUE!
| Mean value | #DIV/0! |  | #VALUE!
(b) | Evaluation of gravimetric test
| Mean volume (Vi) | #VALUE!
| Standard deviation (sr) | #VALUE!
| % A found | #VALUE!
| % CV found | #VALUE!
| Accuracy Limit (A%) | 0.8
| Coefficient Variation Limit (CV%) | 0.3
| Result | #VALUE!
6 (a) | Result of gravimetric test for selected test volume (ul):
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 |  |  | #VALUE!
| X2 |  |  | #VALUE!
| X3 |  |  | #VALUE!
| X4 |  |  | #VALUE!
| X5 |  |  | #VALUE!
| X6 |  |  | #VALUE!
| X7 |  |  | #VALUE!
| X8 |  |  | #VALUE!
| X9 |  |  | #VALUE!
| X10 |  |  | #VALUE!
| Mean value | #DIV/0! |  | #VALUE!
6 (c) | Evaluation of gravimetric test
| Mean volume (Vi) | #VALUE!
| Standard deviation (sd) | #VALUE!
| % A found | #VALUE!
| % CV found | #VALUE!
| Accuracy Limit (A%) | 1.6
| Coefficient Variation Limit (CV%) | 0.6
| Result | #VALUE!
7 (a) | Result of gravimetric test for selected test volume(ul):
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor) | Actual Volume (vol + ML)
| X1 |  |  | #VALUE! | #VALUE!
| X2 |  |  | #VALUE! | #VALUE!
| X3 |  |  | #VALUE! | #VALUE!
| X4 |  |  | #VALUE! | #VALUE!
| X5 |  |  | #VALUE! | #VALUE!
| X6 |  |  | #VALUE! | #VALUE!
| X7 |  |  | #VALUE! | #VALUE!
| X8 |  |  | #VALUE! | #VALUE!
| X9 |  |  | #VALUE! | #VALUE!
| X10 |  |  | #VALUE! | #VALUE!
| Mean value | #DIV/0! |  | #VALUE! | #VALUE!
7 (b) | Mass Loss due to Evaporation
Test Cycle Time (Tsecs)
X11 (mg)
| X11 volume | #VALUE!
| Mass Loss, mg (Mloss) | #VALUE!
7 ( c) | Evaluation of gravimetric test
| Mean volume (Vi) | #VALUE!
| Standard deviation (sd) | #VALUE!
| % A found | #VALUE!
| % CV found | #VALUE!
| Accuracy Limit (A%) | 8
| Coefficient Variation Limit (CV%) | 3
| Result | #VALUE!
8 | Interpretation | Criteria : % A found equal or not more than Accuracy Limit (A%) a% CV found equal or not more than Coefficient Variation Limit (CV%)
Penganalisa |  |  | Penyemak | Pelulus
Tandatangan : Tarikh ujian : Nama : |  |  | Tandatangan : Tarikh semak : Nama : | Tandatangan : Tarikh disahkan : Nama :

## z factor
Temperature | Factor Z | Temperature | Factor Z
(oC) | (mL/g) | (oC) | (mL/g)
15 | 1.002 | 23 | 1.0035
15.5 | 1.002 | 23.5 | 1.0036
16 | 1.0021 | 24 | 1.0038
16.5 | 1.0022 | 24.5 | 1.0039
17 | 1.0023 | 25 | 1.004
17.5 | 1.0024 | 25.5 | 1.0041
18 | 1.0025 | 26 | 1.0043
18.5 | 1.0026 | 26.5 | 1.0044
19 | 1.0027 | 27 | 1.0045
19.5 | 1.0028 | 27.5 | 1.0047
20 | 1.0029 | 28 | 1.0048
20.5 | 1.003 | 28.5 | 1.005
21 | 1.0031 | 29 | 1.0051
21.5 | 1.0032 | 29.5 | 1.0052
22 | 1.0033 | 30 | 1.0054
22.5 | 1.0034

## A&CV NOT IN 6.1
Calculation of maximum permissible errors for volumes not listed in tables.
Nominal volume |  | µL
Selected volume |  | µL
Max error at nominal |  | %
(A or CV)
Max error at selected | #DIV/0! | %
volume
(A or CV)
*Random error (CV) will follow along the systematic error (A).