## custom volume UPM034A
1 | Instrument | Brand Pipettor (code) |  | FINPIPETTE
|  | Serial Number |  | OZ58587
|  | Nominal Volume (ul) |  | 1000
2 | Physical condition | □ Good working condition □ Functional damage : leaking; display not working, button defective, battery defective.
3 | Room condition | Suhu bilik (20 ± 3°C) |  | 22.4
|  | Relative humidity (45 - 80%) |  | 65
4 | Water temperature (˚C) | 22.5
| Balance | METLER TOLEDO XP 205 DR (006-004852)
| Thermometer | THERMO FISHER LOLLIPOP (150736002)
| Z Factor | 1.0034
5 (a) | Result of gravimetric test for selected test volume (ul): |  |  | 1000
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 997.47 |  | 1000.8614
| X2 | 998.07 |  | 1001.4634
| X3 | 996.98 |  | 1000.3697
| X4 | 997.94 |  | 1001.333
| X5 | 998.99 |  | 1002.3866
| X6 | 997.74 |  | 1001.1323
| X7 | 997.66 |  | 1001.052
| X8 | 996.71 |  | 1000.0988
| X9 | 996.52 |  | 999.9082
| X10 | 997.55 |  | 1000.9417
| Mean value | 997.563 |  | 1000.9547
(b) | Evaluation of gravimetric test
| Mean volume (Vi) | 1000.9547
| Standard deviation (sr) | 0.7205
| % A found | 0.095471
| % CV found | 0.071981
| Accuracy Limit (A%) | 0.8
| Coefficient Variation Limit (CV%) | 0.3
| Result | PASS
6 (a) | Result of gravimetric test for selected test volume (ul): |  |  | 500
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 496.3 |  | 497.9874
| X2 | 498.19 |  | 499.8838
| X3 | 497.41 |  | 499.1012
| X4 | 497.48 |  | 499.1714
| X5 | 496.09 |  | 497.7767
| X6 | 496.26 |  | 497.9473
| X7 | 497.82 |  | 499.5126
| X8 | 498.09 |  | 499.7835
| X9 | 497.37 |  | 499.0611
| X10 | 496.83 |  | 498.5192
| Mean value | 497.184 |  | 498.8744
6 (c) | Evaluation of gravimetric test
| Mean volume (Vi) | 498.8744
| Standard deviation (sd) | 0.77493
| % A found | 0.22511
| % CV found | 0.15534
| Accuracy Limit (A%) | 1.6
| Coefficient Variation Limit (CV%) | 0.6
| Result | PASS
7 (a) | Result of gravimetric test for selected test volume(ul): |  |  | 100
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 97.16 |  | 97.4903
| X2 | 96.36 |  | 96.6876
| X3 | 96.64 |  | 96.9686
| X4 | 96.91 |  | 97.2395
| X5 | 97.03 |  | 97.3599
| X6 | 97.05 |  | 97.38
| X7 | 98.92 |  | 99.2563
| X8 | 98.69 |  | 99.0255
| X9 | 96.84 |  | 97.1693
| X10 | 97.1 |  | 97.4301
| Mean value | 97.27 |  | 97.6007
7 ( b) | Evaluation of gravimetric test
| Mean volume (Vi) | 97.6007
| Standard deviation (sd) | 0.84782
| % A found | 2.3993
| % CV found | 0.86866
| Accuracy Limit (A%) | 8
| Coefficient Variation Limit (CV%) | 3
| Result | PASS
8 | Interpretation | Criteria : % A found equal or not more than Accuracy Limit (A%) a% CV found equal or not more than Coefficient Variation Limit (CV%)
Penganalisa |  |  | Penyemak | Pelulus
Tandatangan : Tarikh ujian : Nama : |  |  | Tandatangan : Tarikh semak : Nama : | Tandatangan : Tarikh disahkan : Nama :

## < 50ul UPM034B
1 | Instrument | Brand Pipettor (code) |  | FINPIPETTE
Serial Number
Nominal Volume (ul)
2 | Physical condition | □ Good working condition □ Functional damage : leaking; display not working, button defective, battery defective.
3 | Room condition | Suhu bilik (20 ± 3°C)
Relative humidity (45 - 80%)
4 | Water temperature (˚C)
| Balance | METLER TOLEDO XP 205 DR (006-004852)
| Thermometer | THERMO FISHER LOLLIPOP (150736002)
Z Factor
5 (a) | Result of gravimetric test for selected test volume (ul):
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 |  |  | #VALUE!
| X2 |  |  | #VALUE!
| X3 |  |  | #VALUE!
| X4 |  |  | #VALUE!
| X5 |  |  | #VALUE!
| X6 |  |  | #VALUE!
| X7 |  |  | #VALUE!
| X8 |  |  | #VALUE!
| X9 |  |  | #VALUE!
| X10 |  |  | #VALUE!
| Mean value | #DIV/0! |  | #VALUE!
(b) | Evaluation of gravimetric test
| Mean volume (Vi) | #VALUE!
| Standard deviation (sr) | #VALUE!
| % A found | #VALUE!
| % CV found | #VALUE!
| Accuracy Limit (A%) | 0.8
| Coefficient Variation Limit (CV%) | 0.3
| Result | #VALUE!
6 (a) | Result of gravimetric test for selected test volume (ul):
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 |  |  | #VALUE!
| X2 |  |  | #VALUE!
| X3 |  |  | #VALUE!
| X4 |  |  | #VALUE!
| X5 |  |  | #VALUE!
| X6 |  |  | #VALUE!
| X7 |  |  | #VALUE!
| X8 |  |  | #VALUE!
| X9 |  |  | #VALUE!
| X10 |  |  | #VALUE!
| Mean value | #DIV/0! |  | #VALUE!
6 (c) | Evaluation of gravimetric test
| Mean volume (Vi) | #VALUE!
| Standard deviation (sd) | #VALUE!
| % A found | #VALUE!
| % CV found | #VALUE!
| Accuracy Limit (A%) | 1.6
| Coefficient Variation Limit (CV%) | 0.6
| Result | #VALUE!
7 (a) | Result of gravimetric test for selected test volume(ul):
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor) | Actual Volume (vol + ML)
| X1 |  |  | #VALUE! | #VALUE!
| X2 |  |  | #VALUE! | #VALUE!
| X3 |  |  | #VALUE! | #VALUE!
| X4 |  |  | #VALUE! | #VALUE!
| X5 |  |  | #VALUE! | #VALUE!
| X6 |  |  | #VALUE! | #VALUE!
| X7 |  |  | #VALUE! | #VALUE!
| X8 |  |  | #VALUE! | #VALUE!
| X9 |  |  | #VALUE! | #VALUE!
| X10 |  |  | #VALUE! | #VALUE!
| Mean value | #DIV/0! |  | #VALUE! | #VALUE!
7 (b) | Mass Loss due to Evaporation
Test Cycle Time (Tsecs)
X11 (mg)
| X11 volume | #VALUE!
| Mass Loss, mg (Mloss) | #VALUE!
7 ( c) | Evaluation of gravimetric test
| Mean volume (Vi) | #VALUE!
| Standard deviation (sd) | #VALUE!
| % A found | #VALUE!
| % CV found | #VALUE!
| Accuracy Limit (A%) | 8
| Coefficient Variation Limit (CV%) | 3
| Result | #VALUE!
8 | Interpretation | Criteria : % A found equal or not more than Accuracy Limit (A%) a% CV found equal or not more than Coefficient Variation Limit (CV%)
Penganalisa |  |  | Penyemak | Pelulus
Tandatangan : Tarikh ujian : Nama : |  |  | Tandatangan : Tarikh semak : Nama : | Tandatangan : Tarikh disahkan : Nama :

## z factor
Temperature | Factor Z | Temperature | Factor Z
(oC) | (mL/g) | (oC) | (mL/g)
15 | 1.002 | 23 | 1.0035
15.5 | 1.002 | 23.5 | 1.0036
16 | 1.0021 | 24 | 1.0038
16.5 | 1.0022 | 24.5 | 1.0039
17 | 1.0023 | 25 | 1.004
17.5 | 1.0024 | 25.5 | 1.0041
18 | 1.0025 | 26 | 1.0043
18.5 | 1.0026 | 26.5 | 1.0044
19 | 1.0027 | 27 | 1.0045
19.5 | 1.0028 | 27.5 | 1.0047
20 | 1.0029 | 28 | 1.0048
20.5 | 1.003 | 28.5 | 1.005
21 | 1.0031 | 29 | 1.0051
21.5 | 1.0032 | 29.5 | 1.0052
22 | 1.0033 | 30 | 1.0054
22.5 | 1.0034

## A&CV NOT IN 6.1
Calculation of maximum permissible errors for volumes not listed in tables.
Nominal volume |  | µL
Selected volume |  | µL
Max error at nominal |  | %
(A or CV)
Max error at selected | #DIV/0! | %
volume
(A or CV)
*Random error (CV) will follow along the systematic error (A).