## 150736002
Additional Document/ Worksheet
GENERATION OF LINEARITY EQUATION FOR ACTUAL TEMPERATURE READING
ID of Test Thermometer |  | 150736002
Certification Number |  | MSRP24090084 | Calibration Due Date |  | 2026-09-04
| Test Thermometer Temperature Reading (°C) | Correction | Actual Temperature Reading (°C)
| -0.5 | 0.5 | 0
| 19.8 | 0.2 | 20
| 49.5 | 0.5 | 50
| 79.4 | 0.6 | 80
| 99.4 | 0.6 | 100
Calculation:
Linearity equation for correction |  | Y = | 1.0025 | X + | 0.3564
X = Test Thermometer Reading; Y = Actual Temperature Reading
Test Thermometer Temperature Reading (°C) |  | 21.7
Actual Temperature Reading (°C) |  | 22.1106

## temperature list
temperature list
Brand Pipettor (code) | Serial Number | Nominal Volume (μl)
FINPIPETTE® F2 | NZ47544 | 5000
FINPIPETTE® F2 | NZ02006 | 5000
FINPIPETTE® F2 | OZ58589 | 1000
FINPIPETTE® F2 | PZ06943 | 100

## 111668831
Additional Document/ Worksheet
GENERATION OF LINEARITY EQUATION FOR ACTUAL TEMPERATURE READING
ID of Test Thermometer |  | 111668831
Certification Number |  | AI.T2209.S0996 | Calibration Due Date |  | 2024-09-07
| Test Thermometer Temperature Reading (°C) | Correction | Actual Temperature Reading (°C)
| 20.2 | -0.2 | 20
| 25.3 | -0.3 | 25
| 30.6 | -0.6 | 30
Calculation:
Linearity equation for correction |  | Y = |  | X +
X = Test Thermometer Reading; Y = Actual Temperature Reading
Test Thermometer Temperature Reading (°C)
Actual Temperature Reading (°C) |  | 0