## custom volume UPM034A
1 | Instrument | Brand Pipettor (code) |  | FINPIPETTE
|  | Serial Number |  | NZ47544
|  | Nominal Volume (ul) |  | 5000
2 | Physical condition | □ Good working condition □ Functional damage : leaking; display not working, button defective, battery defective.
3 | Room condition | Suhu bilik (20 ± 3°C) |  | 21.6
|  | Relative humidity (45 - 80%) |  | 65.8
4 | Water temperature (˚C) | 22.0104
| Balance | METLER TOLEDO XP 205 DR (006-004852)
| Thermometer | THERMO FISHER LOLLIPOP (150736002)
| Z Factor | 1.0033
5 (a) | Result of gravimetric test for selected test volume (ul): |  |  | 5000
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 4950.14 |  | 4966.4755
| X2 | 4965.19 |  | 4981.5751
| X3 | 4955.73 |  | 4972.0839
| X4 | 4936.72 |  | 4953.0112
| X5 | 4965.77 |  | 4982.157
| X6 | 4957.77 |  | 4974.1306
| X7 | 4952.04 |  | 4968.3817
| X8 | 4958.14 |  | 4974.5019
| X9 | 4954.69 |  | 4971.0405
| X10 | 4958.41 |  | 4974.7728
| Mean value | 4955.46 |  | 4971.813
(b) | Evaluation of gravimetric test
| Mean volume (Vi) | 4971.813
| Standard deviation (sr) | 8.2872
| % A found | 0.56374
| % CV found | 0.16668
| Accuracy Limit (A%) | 0.8
| Coefficient Variation Limit (CV%) | 0.3
| Result | PASS
6 (a) | Result of gravimetric test for selected test volume (ul): |  |  | 2500
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 2492.76 |  | 2500.9861
| X2 | 2474.77 |  | 2482.9367
| X3 | 2479.32 |  | 2487.5018
| X4 | 2478.65 |  | 2486.8295
| X5 | 2483.6 |  | 2491.7959
| X6 | 2486.31 |  | 2494.5148
| X7 | 2484.88 |  | 2493.0801
| X8 | 2479.34 |  | 2487.5218
| X9 | 2472.71 |  | 2480.8699
| X10 | 2484.86 |  | 2493.06
| Mean value | 2481.72 |  | 2489.9097
6 (c) | Evaluation of gravimetric test
| Mean volume (Vi) | 2489.9097
| Standard deviation (sd) | 5.9512
| % A found | 0.40361
| % CV found | 0.23901
| Accuracy Limit (A%) | 1.6
| Coefficient Variation Limit (CV%) | 0.6
| Result | PASS
7 (a) | Result of gravimetric test for selected test volume(ul): |  |  | 500
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 490.21 |  | 491.8277
| X2 | 490.81 |  | 492.4297
| X3 | 497.88 |  | 499.523
| X4 | 494.24 |  | 495.871
| X5 | 490.59 |  | 492.2089
| X6 | 486.96 |  | 488.567
| X7 | 488.83 |  | 490.4431
| X8 | 491.65 |  | 493.2724
| X9 | 487.95 |  | 489.5602
| X10 | 488.91 |  | 490.5234
| Mean value | 490.803 |  | 492.4226
7 ( b) | Evaluation of gravimetric test
| Mean volume (Vi) | 492.4226
| Standard deviation (sd) | 3.232
| % A found | 1.5155
| % CV found | 0.65635
| Accuracy Limit (A%) | 8
| Coefficient Variation Limit (CV%) | 3
| Result | PASS
8 | Interpretation | Criteria : % A found equal or not more than Accuracy Limit (A%) a% CV found equal or not more than Coefficient Variation Limit (CV%)
Penganalisa |  |  | Penyemak | Pelulus
Tandatangan : Tarikh ujian : Nama : |  |  | Tandatangan : Tarikh semak : Nama : | Tandatangan : Tarikh disahkan : Nama :

## < 50ul UPM034B
1 | Instrument | Brand Pipettor (code) |  | FINPIPETTE
|  | Serial Number |  | FJ04697
|  | Nominal Volume (ul) |  | 100
2 | Physical condition | □ Good working condition □ Functional damage : leaking; display not working, button defective, battery defective.
3 | Room condition | Suhu bilik (20 ± 3°C) |  | 21.6
|  | Relative humidity (45 - 80%) |  | 65.8
4 | Water temperature (˚C) | 22.0104
| Balance | METLER TOLEDO XP 205 DR (006-004852)
| Thermometer | THERMO FISHER LOLLIPOP (150736002)
| Z Factor | 1.0033
5 (a) | Result of gravimetric test for selected test volume (ul):
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 |  |  | 0
| X2 |  |  | 0
| X3 |  |  | 0
| X4 |  |  | 0
| X5 |  |  | 0
| X6 |  |  | 0
| X7 |  |  | 0
| X8 |  |  | 0
| X9 |  |  | 0
| X10 |  |  | 0
| Mean value | #DIV/0! |  | 0
(b) | Evaluation of gravimetric test
| Mean volume (Vi) | 0
| Standard deviation (sr) | 0
| % A found | #DIV/0!
| % CV found | #DIV/0!
| Accuracy Limit (A%) | 0.8
| Coefficient Variation Limit (CV%) | 0.3
| Result | #DIV/0!
6 (a) | Result of gravimetric test for selected test volume (ul):
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 |  |  | 0
| X2 |  |  | 0
| X3 |  |  | 0
| X4 |  |  | 0
| X5 |  |  | 0
| X6 |  |  | 0
| X7 |  |  | 0
| X8 |  |  | 0
| X9 |  |  | 0
| X10 |  |  | 0
| Mean value | #DIV/0! |  | 0
6 (c) | Evaluation of gravimetric test
| Mean volume (Vi) | 0
| Standard deviation (sd) | 0
| % A found | #DIV/0!
| % CV found | #DIV/0!
| Accuracy Limit (A%) | 1.6
| Coefficient Variation Limit (CV%) | 0.6
| Result | #DIV/0!
7 (a) | Result of gravimetric test for selected test volume(ul): |  |  | 10
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor) | Actual Volume (vol + ML)
| X1 | 9.23 |  | 9.2605 | 9.2946
| X2 | 9.56 |  | 9.5915 | 9.6257
| X3 | 9.65 |  | 9.6818 | 9.716
| X4 | 9.49 |  | 9.5213 | 9.5554
| X5 | 9.48 |  | 9.5113 | 9.5454
| X6 | 9.66 |  | 9.6919 | 9.726
| X7 | 9.11 |  | 9.1401 | 9.1742
| X8 | 9.14 |  | 9.1702 | 9.2043
| X9 | 9.14 |  | 9.1702 | 9.2043
| X10 | 9.16 |  | 9.1902 | 9.2243
| Mean value | 9.362 |  | 9.3929 | 9.427
7 (b) | Mass Loss due to Evaporation
| Test Cycle Time (Tsecs) | 60
| X11 (mg) | 8.82
| X11 volume | 8.8491
| Mass Loss, mg (Mloss) | 0.034112
7 ( c) | Evaluation of gravimetric test
| Mean volume (Vi) | 9.427
| Standard deviation (sd) | 0.22721
| % A found | 5.7299
| % CV found | 2.4102
| Accuracy Limit (A%) | 8
| Coefficient Variation Limit (CV%) | 3
| Result | PASS
8 | Interpretation | Criteria : % A found equal or not more than Accuracy Limit (A%) a% CV found equal or not more than Coefficient Variation Limit (CV%)
Penganalisa |  |  | Penyemak | Pelulus
Tandatangan : Tarikh ujian : Nama : |  |  | Tandatangan : Tarikh semak : Nama : | Tandatangan : Tarikh disahkan : Nama :

## z factor
Temperature | Factor Z | Temperature | Factor Z
(oC) | (mL/g) | (oC) | (mL/g)
15 | 1.002 | 23 | 1.0035
15.5 | 1.002 | 23.5 | 1.0036
16 | 1.0021 | 24 | 1.0038
16.5 | 1.0022 | 24.5 | 1.0039
17 | 1.0023 | 25 | 1.004
17.5 | 1.0024 | 25.5 | 1.0041
18 | 1.0025 | 26 | 1.0043
18.5 | 1.0026 | 26.5 | 1.0044
19 | 1.0027 | 27 | 1.0045
19.5 | 1.0028 | 27.5 | 1.0047
20 | 1.0029 | 28 | 1.0048
20.5 | 1.003 | 28.5 | 1.005
21 | 1.0031 | 29 | 1.0051
21.5 | 1.0032 | 29.5 | 1.0052
22 | 1.0033 | 30 | 1.0054
22.5 | 1.0034

## A&CV NOT IN 6.1
Calculation of maximum permissible errors for volumes not listed in tables.
Nominal volume |  | µL
Selected volume |  | µL
Max error at nominal |  | %
(A or CV)
Max error at selected | #DIV/0! | %
volume
(A or CV)
*Random error (CV) will follow along the systematic error (A).