## custom volume UPM034A
1 | Instrument | Brand Pipettor (code)
Serial Number
Nominal Volume (ul)
2 | Physical condition | □ Good working condition □ Functional damage : leaking; display not working, button defective, battery defective.
3 | Room condition | Suhu bilik (20 ± 3°C) |  | 21.6
|  | Relative humidity (45 - 80%) |  | 62.9
4 | Water temperature (˚C) | 21.5078
Balance
Thermometer
| Z Factor | 1.0032
5 (a) | Result of gravimetric test for selected test volume (ul): |  |  | 100
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 100 |  | 100.32
| X2 | 100.51 |  | 100.8316
| X3 | 100.2 |  | 100.5206
| X4 | 100 |  | 100.32
| X5 | 100 |  | 100.32
| X6 | 100.7 |  | 101.0222
| X7 | 100 |  | 100.32
| X8 | 100.58 |  | 100.9019
| X9 | 100.2 |  | 100.5206
| X10 | 100.01 |  | 100.33
| Mean value | 100.22 |  | 100.5407
(b) | Evaluation of gravimetric test
| Mean volume (Vi) | 100.5407
| Standard deviation (sr) | 0.27628
| % A found | 0.5407
| % CV found | 0.27479
| Accuracy Limit (A%) | 0.8
| Coefficient Variation Limit (CV%) | 0.3
| Result | PASS
6 (a) | Result of gravimetric test for selected test volume (ul): |  |  | 500
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 2509 |  | 2517.0288
| X2 | 2504 |  | 2512.0128
| X3 | 2501 |  | 2509.0032
| X4 | 2499 |  | 2506.9968
| X5 | 2505 |  | 2513.016
| X6 | 2500 |  | 2508
| X7 | 2504 |  | 2512.0128
| X8 | 2499 |  | 2506.9968
| X9 | 2505 |  | 2513.016
| X10 | 2505 |  | 2513.016
| Mean value | 2503.1 |  | 2511.1099
6 (c) | Evaluation of gravimetric test
| Mean volume (Vi) | 2511.1099
| Standard deviation (sd) | 3.2576
| % A found | 402.222
| % CV found | 0.12973
| Accuracy Limit (A%) | 1.6
| Coefficient Variation Limit (CV%) | 0.6
| Result | FAIL
7 (a) | Result of gravimetric test for selected test volume(ul): |  |  | 500
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 506 |  | 507.6192
| X2 | 507 |  | 508.6224
| X3 | 510 |  | 511.632
| X4 | 509 |  | 510.6288
| X5 | 512 |  | 513.6384
| X6 | 507 |  | 508.6224
| X7 | 503 |  | 504.6096
| X8 | 509 |  | 510.6288
| X9 | 507 |  | 508.6224
| X10 | 504 |  | 505.6128
| Mean value | 507 |  | 509.0237
7 ( b) | Evaluation of gravimetric test
| Mean volume (Vi) | 509.0237
| Standard deviation (sd) | 2.7249
| % A found | 1.8047
| % CV found | 0.53532
| Accuracy Limit (A%) | 8
| Coefficient Variation Limit (CV%) | 3
| Result | PASS
8 | Interpretation | Criteria : % A found equal or not more than Accuracy Limit (A%) a% CV found equal or not more than Coefficient Variation Limit (CV%)
Penganalisa |  |  | Penyemak | Pelulus
Tandatangan : Tarikh ujian : Nama : |  |  | Tandatangan : Tarikh semak : Nama : | Tandatangan : Tarikh disahkan : Nama :

## < 50ul UPM034B
1 | Instrument | Brand Pipettor (code)
Serial Number
Nominal Volume (ul)
2 | Physical condition | □ Good working condition □ Functional damage : leaking; display not working, button defective, battery defective.
3 | Room condition | Suhu bilik (20 ± 3°C)
Relative humidity (45 - 80%)
4 | Water temperature (˚C)
Balance
Thermometer
Z Factor
5 (a) | Result of gravimetric test for selected test volume (ul):
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 |  |  | #VALUE!
| X2 |  |  | #VALUE!
| X3 |  |  | #VALUE!
| X4 |  |  | #VALUE!
| X5 |  |  | #VALUE!
| X6 |  |  | #VALUE!
| X7 |  |  | #VALUE!
| X8 |  |  | #VALUE!
| X9 |  |  | #VALUE!
| X10 |  |  | #VALUE!
| Mean value | #DIV/0! |  | #VALUE!
(b) | Evaluation of gravimetric test
| Mean volume (Vi) | #VALUE!
| Standard deviation (sr) | #VALUE!
| % A found | #VALUE!
| % CV found | #VALUE!
Accuracy Limit (A%)
Coefficient Variation Limit (CV%)
| Result | #VALUE!
6 (a) | Result of gravimetric test for selected test volume (ul):
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 |  |  | #VALUE!
| X2 |  |  | #VALUE!
| X3 |  |  | #VALUE!
| X4 |  |  | #VALUE!
| X5 |  |  | #VALUE!
| X6 |  |  | #VALUE!
| X7 |  |  | #VALUE!
| X8 |  |  | #VALUE!
| X9 |  |  | #VALUE!
| X10 |  |  | #VALUE!
| Mean value | #DIV/0! |  | #VALUE!
6 (c) | Evaluation of gravimetric test
| Mean volume (Vi) | #VALUE!
| Standard deviation (sd) | #VALUE!
| % A found | #VALUE!
| % CV found | #VALUE!
Accuracy Limit (A%)
Coefficient Variation Limit (CV%)
| Result | #VALUE!
7 (a) | Result of gravimetric test for selected test volume(ul):
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor) | Actual Volume (vol + ML)
| X1 |  |  | #VALUE! | #VALUE!
| X2 |  |  | #VALUE! | #VALUE!
| X3 |  |  | #VALUE! | #VALUE!
| X4 |  |  | #VALUE! | #VALUE!
| X5 |  |  | #VALUE! | #VALUE!
| X6 |  |  | #VALUE! | #VALUE!
| X7 |  |  | #VALUE! | #VALUE!
| X8 |  |  | #VALUE! | #VALUE!
| X9 |  |  | #VALUE! | #VALUE!
| X10 |  |  | #VALUE! | #VALUE!
| Mean value | #DIV/0! |  | #VALUE! | #VALUE!
7 (b) | Mass Loss due to Evaporation
Test Cycle Time (Tsecs)
X11 (mg)
| X11 volume | #VALUE!
| Mass Loss, mg (Mloss) | #VALUE!
7 ( c) | Evaluation of gravimetric test
| Mean volume (Vi) | #VALUE!
| Standard deviation (sd) | #VALUE!
| % A found | #VALUE!
| % CV found | #VALUE!
Accuracy Limit (A%)
Coefficient Variation Limit (CV%)
| Result | #VALUE!
8 | Interpretation | Criteria : % A found equal or not more than Accuracy Limit (A%) a% CV found equal or not more than Coefficient Variation Limit (CV%)
Penganalisa |  |  | Penyemak | Pelulus
Tandatangan : Tarikh ujian : Nama : |  |  | Tandatangan : Tarikh semak : Nama : | Tandatangan : Tarikh disahkan : Nama :

## z factor
Temperature | Factor Z | Temperature | Factor Z
(oC) | (mL/g) | (oC) | (mL/g)
15 | 1.002 | 23 | 1.0035
15.5 | 1.002 | 23.5 | 1.0036
16 | 1.0021 | 24 | 1.0038
16.5 | 1.0022 | 24.5 | 1.0039
17 | 1.0023 | 25 | 1.004
17.5 | 1.0024 | 25.5 | 1.0041
18 | 1.0025 | 26 | 1.0043
18.5 | 1.0026 | 26.5 | 1.0044
19 | 1.0027 | 27 | 1.0045
19.5 | 1.0028 | 27.5 | 1.0047
20 | 1.0029 | 28 | 1.0048
20.5 | 1.003 | 28.5 | 1.005
21 | 1.0031 | 29 | 1.0051
21.5 | 1.0032 | 29.5 | 1.0052
22 | 1.0033 | 30 | 1.0054
22.5 | 1.0034

## A&CV NOT IN 6.1
Calculation of maximum permissible errors for volumes not listed in tables.
Nominal volume |  | µL
Selected volume |  | µL
Max error at nominal |  | %
(A or CV)
Max error at selected | #DIV/0! | %
volume
(A or CV)
*Random error (CV) will follow along the systematic error (A).