## custom volume UPM034A
1 | Instrument | Brand Pipettor (code) |  | FINPIPETTE
|  | Serial Number |  | NZ47544
|  | Nominal Volume (ul) |  | 5000
2 | Physical condition | □ Good working condition □ Functional damage : leaking; display not working, button defective, battery defective.
3 | Room condition | Suhu bilik (20 ± 3°C) |  | 21.3
|  | Relative humidity (45 - 80%) |  | 61.9
4 | Water temperature (˚C) | 22.2109
| Balance | METLER TOLEDO XP 205 DR (006-004852)
| Thermometer | THERMO FISHER LOLLIPOP (150736002)
| Z Factor | 1.0033
5 (a) | Result of gravimetric test for selected test volume (ul): |  |  | 5000
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 5009.7 |  | 5026.232
| X2 | 5003.3 |  | 5019.8109
| X3 | 5003.8 |  | 5020.3125
| X4 | 5001.8 |  | 5018.3059
| X5 | 5003.9 |  | 5020.4129
| X6 | 5000.5 |  | 5017.0017
| X7 | 5005.3 |  | 5021.8175
| X8 | 5004.6 |  | 5021.1152
| X9 | 5009.4 |  | 5025.931
| X10 | 5007.8 |  | 5024.3257
| Mean value | 5005.01 |  | 5021.5265
(b) | Evaluation of gravimetric test
| Mean volume (Vi) | 5021.5265
| Standard deviation (sr) | 3.0936
| % A found | 0.43053
| % CV found | 0.061607
| Accuracy Limit (A%) | 0.8
| Coefficient Variation Limit (CV%) | 0.3
| Result | PASS
6 (a) | Result of gravimetric test for selected test volume (ul): |  |  | 2500
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 2501.7 |  | 2509.9556
| X2 | 2508.6 |  | 2516.8784
| X3 | 2509.4 |  | 2517.681
| X4 | 2504.3 |  | 2512.5642
| X5 | 2503.5 |  | 2511.7616
| X6 | 2508.6 |  | 2516.8784
| X7 | 2508.6 |  | 2516.8784
| X8 | 2504.7 |  | 2512.9655
| X9 | 2503.5 |  | 2511.7616
| X10 | 2503 |  | 2511.2599
| Mean value | 2505.59 |  | 2513.8584
6 (c) | Evaluation of gravimetric test
| Mean volume (Vi) | 2513.8584
| Standard deviation (sd) | 2.891
| % A found | 0.55434
| % CV found | 0.115
| Accuracy Limit (A%) | 1.6
| Coefficient Variation Limit (CV%) | 0.6
| Result | PASS
7 (a) | Result of gravimetric test for selected test volume(ul): |  |  | 500
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 505.27 |  | 506.9374
| X2 | 504.79 |  | 506.4558
| X3 | 502.69 |  | 504.3489
| X4 | 505.41 |  | 507.0779
| X5 | 508.31 |  | 509.9874
| X6 | 508.31 |  | 509.9874
| X7 | 505.92 |  | 507.5895
| X8 | 509.43 |  | 511.1111
| X9 | 505.42 |  | 507.0879
| X10 | 504.94 |  | 506.6063
| Mean value | 507 |  | 507.719
7 ( b) | Evaluation of gravimetric test
| Mean volume (Vi) | 507.719
| Standard deviation (sd) | 2.0392
| % A found | 1.5438
| % CV found | 0.40165
| Accuracy Limit (A%) | 8
| Coefficient Variation Limit (CV%) | 3
| Result | PASS
8 | Interpretation | Criteria : % A found equal or not more than Accuracy Limit (A%) a% CV found equal or not more than Coefficient Variation Limit (CV%)
Penganalisa |  |  | Penyemak | Pelulus
Tandatangan : Tarikh ujian : Nama : |  |  | Tandatangan : Tarikh semak : Nama : | Tandatangan : Tarikh disahkan : Nama :

## < 50ul UPM034B
1 | Instrument | Brand Pipettor (code) |  | FINPIPETTE
|  | Serial Number |  | NZ47544
|  | Nominal Volume (ul) |  | 5000
2 | Physical condition | □ Good working condition □ Functional damage : leaking; display not working, button defective, battery defective.
3 | Room condition | Suhu bilik (20 ± 3°C) |  | 21.3
|  | Relative humidity (45 - 80%) |  | 61.9
4 | Water temperature (˚C) | 22.2095
| Balance | METLER TOLEDO XP 205 DR (006-004852)
| Thermometer | THERMO FISHER LOLLIPOP (150736002)
| Z Factor | 1.0033
5 (a) | Result of gravimetric test for selected test volume (ul): |  |  | 5000
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 5009.7 |  | 5026.232
| X2 | 5003.3 |  | 5019.8109
| X3 | 5003.8 |  | 5020.3125
| X4 | 5001.8 |  | 5018.3059
| X5 | 5003.9 |  | 5020.4129 |  | M
| X6 | 5000.5 |  | 5017.0017
| X7 | 5005.3 |  | 5021.8175
| X8 | 5004.6 |  | 5021.1152
| X9 | 5009.4 |  | 5025.931
| X10 | 5007.8 |  | 5024.3257
| Mean value | 5005.01 |  | 5021.5265
(b) | Evaluation of gravimetric test
| Mean volume (Vi) | 5021.5265
| Standard deviation (sr) | 3.0936
| % A found | 0.43053
| % CV found | 0.061607
| Accuracy Limit (A%) | 0.8
| Coefficient Variation Limit (CV%) | 0.3
| Result | PASS
6 (a) | Result of gravimetric test for selected test volume (ul): |  |  | 2500
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor)
| X1 | 2501.7 |  | 2509.9556
| X2 | 2508.6 |  | 2516.8784
| X3 | 2509.4 |  | 2517.681
| X4 | 2504.3 |  | 2512.5642
| X5 | 2503.5 |  | 2511.7616
| X6 | 2508.6 |  | 2516.8784
| X7 | 2508.6 |  | 2516.8784
| X8 | 2504.7 |  | 2512.9655
| X9 | 2503.5 |  | 2511.7616
| X10 | 2503 |  | 2511.2599
| Mean value | 2505.59 |  | 2513.8584
6 (c) | Evaluation of gravimetric test
| Mean volume (Vi) | 2513.8584
| Standard deviation (sd) | 2.891
| % A found | 0.55434
| % CV found | 0.115
| Accuracy Limit (A%) | 1.6
| Coefficient Variation Limit (CV%) | 0.6
| Result | PASS
7 (a) | Result of gravimetric test for selected test volume(ul): |  |  | 500
| Weighing No. | Weight (mg) |  | Converted Volume (mass * Z factor) | Actual Volume (vol + ML)
| X1 | 505.27 |  | 506.9374 | 557.8759
| X2 | 504.79 |  | 506.4558 | 557.3944
| X3 | 507 |  | 508.6731 | 559.6116
| X4 | 507.71 |  | 509.3854 | 560.324
| X5 | 508.69 |  | 510.3687 | 561.3072
| X6 | 503.34 |  | 505.001 | 555.9396
| X7 | 506.26 |  | 507.9307 | 558.8692
| X8 | 505.83 |  | 507.4992 | 558.4378
| X9 | 507.32 |  | 508.9942 | 559.9327
| X10 | 507.71 |  | 509.3854 | 560.324
| Mean value | 506.392 |  | 508.0631 | 559.0016
7 (b) | Mass Loss due to Evaporation
Test Cycle Time (Tsecs)
X11 (mg)
| X11 volume | 0
| Mass Loss, mg (Mloss) | 50.9385
7 ( c) | Evaluation of gravimetric test
| Mean volume (Vi) | 559.0016
| Standard deviation (sd) | 1.6198
| % A found | 11.8003
| % CV found | 0.28977
| Accuracy Limit (A%) | 8
| Coefficient Variation Limit (CV%) | 3
| Result | FAIL
8 | Interpretation | Criteria : % A found equal or not more than Accuracy Limit (A%) a% CV found equal or not more than Coefficient Variation Limit (CV%)
Penganalisa |  |  | Penyemak | Pelulus
Tandatangan : Tarikh ujian : Nama : |  |  | Tandatangan : Tarikh semak : Nama : | Tandatangan : Tarikh disahkan : Nama :

## z factor
Temperature | Factor Z | Temperature | Factor Z
(oC) | (mL/g) | (oC) | (mL/g)
15 | 1.002 | 23 | 1.0035
15.5 | 1.002 | 23.5 | 1.0036
16 | 1.0021 | 24 | 1.0038
16.5 | 1.0022 | 24.5 | 1.0039
17 | 1.0023 | 25 | 1.004
17.5 | 1.0024 | 25.5 | 1.0041
18 | 1.0025 | 26 | 1.0043
18.5 | 1.0026 | 26.5 | 1.0044
19 | 1.0027 | 27 | 1.0045
19.5 | 1.0028 | 27.5 | 1.0047
20 | 1.0029 | 28 | 1.0048
20.5 | 1.003 | 28.5 | 1.005
21 | 1.0031 | 29 | 1.0051
21.5 | 1.0032 | 29.5 | 1.0052
22 | 1.0033 | 30 | 1.0054
22.5 | 1.0034

## A&CV NOT IN 6.1
Calculation of maximum permissible errors for volumes not listed in tables.
Nominal volume |  | µL
Selected volume |  | µL
Max error at nominal |  | %
(A or CV)
Max error at selected | #DIV/0! | %
volume
(A or CV)
*Random error (CV) will follow along the systematic error (A).